## Slide 1
Arbitrary Waveform Generator
DTRI EPD-Oct 2024 –Team 2

## Slide 2
Sl No | Parameter | Specifications | Remarks
1 | Max Frequency | 100 MHz | 
2 | Bandwidth | 10 MHz | Changed from 40 to 10 MHz after Review No.3
3 | Max Output Voltage | <whatever DAC supports is Okay> | 
4 | Sampling rate | <whatever DAC supports is Okay> | 
5 | Resolution | 12-bit DAC | 
6 | Memory | <whatever STM supports is Okay> | 
7 | Waveform repetition support | Yes | 
8 | Standard Waveforms | Sine,Cosine,Square(Variable Duty Cycle),Noise | 
9 | Userdefined waveforms | Accept *.txt (Integer,Hex,Float values) | 
10 | Channel | Single ( Option to go dual ) | 
11 | Clock Output | Yes | 
12 | UI | User defined - waveform type, frequency,bandwidth,duty cycle,output voltage level | 
13 | Impedance | 50 Ohms | 
14 | Interface | USB to PC | 
15 | Dimensions | LxWxH ( 4inX2inX2in) | 
16 | Weight | <500gms | 
17 | Connectors | 3BNC | 
18 | Power | USB or Adaptor | 
Specifications

## Slide 3
Use GNU Octave to generate waves and check their outputs – Text files generated during python
Based on the memory capacity of STM you decide whether to retain floating point inputs or not
Thoroughly check and understand DAC requirements and specs, DAC sampling rate, frequency, bandwidth etc
Check what will be the full-scale output of DAC. This is a very important feature
The noise requirement is AWGN – Arbitrary White Gaussian Noise – Study more on noise generation, check whether DAC support Noise generation, if so get the details
Memory capacity to be finalized based on the input – considering floating point input
Give the clock frequency output in the range of 1 MHz to 10 MHz
Bandwidth can be kept as 10MHz and for frequency values of less than 10 MHz take the bandwidth as 10% of the required output frequency. Bandwidth is a requirement for both Arb waves and for noise – Specification changed accordingly
9. Study the specs of STM 32 446 series
10. Elaborate the block diagram with finer details.
11. Are you going to use USB as the the input for passing .txt files to STM ?
12. Study the frequency domain of white noise
13. Roles and responsibilities and timeline has to be elaborated and mention the details based on the project development headings.
Project review No.3 dt.01.01.2025
Take Aways / Homework points :

## Slide 4
User inputs in the arb directly can be taken from the touch screen for standard wave forms
Waves can be in 3 phases : Equation based, Storage based, Arb (Blind values ie no reference )
A wave digitally is represented as discrete values of time ( sampling rate ) and discrete values of amplitude ( DAC )
Clock frequency output to user can be fixed. It need not be variable
4. Use .txt files as standard format for taking inputs
5. Duty cycle is an input that is required for generation of square waves. This input has to be taken from the user
6. Two channels are necessary as mentioned in the spec sheet
7. All the calculations for wave forms are wrt to time domain and not wrt frequency domain
Project review No.3 dt.01.01.2025
Inputs / Clarifications :

## Slide 5
Check Digilant website
The max scale reading of DAC of 3.6 Volts is okay
What will the clock frequency for output to the customer
Prepare a more granular block diagram
Have the display within 2” and what will the display in the screen ( Real time wave form generation is not necessary on the screen, if the screen displays waveform characteristics that is sufficient )
Check whether battery is necessary, if so what spec battery will be used and it should be a rechargeable battery
What should be the input power and how the power input will be given. Prepare a power budget for your equipment
Project review No.4 dt.13.01.2025
Take Aways / Homework points :

## Slide 6
Project review No.5 dt.15.01.2025
Take Aways / Homework points :
Revision history of previous revisions need to be properly edited ( Resp : PC ) – A/O
Key features in the documentation file should include technical details ( Resp : VK ) – A/O
Prepare a detailed block diagram with the understanding of component selections as on date (Resp : PC) – A/O
Prepare a flow chart for usage of touch screen and create screens with proper color scheme and UI ( Resp : VK ) – A/O
Check for data sheets of AWG of various manufacturers and prepare a similar data sheet for your equipment as well ( Resp : PC ) - Pending
Use a terminal application and transfer a .txt file to the microcontroller ( Resp : SV ) - Pending
Prepare a software architecture for your project ( Resp : PC ) - Pending
Finalize the mechanical dimensions of the AWG ( Depends of Screen selection ) ( Resp : VK ) –( 6” X 5” X 4” )
Select the appropriate touch screen ( You may check DWIN website and also download the suitable SDK ) ( Resp : VK ) - WIP
Update the Roles and responsibilities ( Resp : PC ) – A/O
Check a demo frequency multiplication POC using ST446 microcontroller / Use of External DAC ( Resp : SV ) - WIP
Along with Mouser you can also check for parts availability in Digikey website as well
Have multiple options like file transfer from PC to MC through USD, Use of SD card and entering custom options into the AWG

## Slide 7
Project review No.6 dt.22.01.2025
Sl No | Review Point | Resp | Status | Remarks
1 | Update Key features and specifications. One Example of specifications of Fluke 80 / 81 series can be referred and prepare a similar specification table for your product | VK | Pending | 
2 | Check and confirm whether the AWG has to be equipped with 2 channels or 1 channel is sufficient | VK | A/O | 2 Channels needed
3 | Re-do R&R, project timeline and key features of the AWG | PC | A/O | 
4 | Decide on the number of voltage regulators that will be required. Use LDO type voltage regulators | VK | A/O | 
5 | Details of DC Adaptor ( Whether mounted to board or external, power outputs etc) to be studied and finalized | VK | A/O | 
6 | Do some more research of MMC to understand how the SD card can transfer data to the MCU | SV | A/O | 
7 | Finalize the type of USB that will be used for transfer of data from PC to MCU | VK | A/O | 
8 | Finalize the MCU that will be required to this AWG | SV | A/O | 
9 | Study DACs max current rating and arrive at the resistance rating to get the required final output voltage and check and confirm the necessity of an output amplifier | VK | A/O | 
10 | In the block diagram – mentioned the Clock Input and Output sections separately and the user clock output should be fed from the GPIO pin and not from Clock Output as mentioned in the block diagram | SV | A/O | Changes in the block diagram
11 | The current flow at the output of DAC needs to be corrected | SV | A/O | 
12 | VDD and VDDA should be de-coupled the avoid noise infusion into VDDA. If possible VREF also to be decoupled | PC | A/O | 
13 | The layout of traces from MCU to DAC should be such that the lengths of traces should be the same | PC | Input | 
14 | Try to keep the thickness of the AWG to around 2” | VK | Input | 

## Slide 8
Development of Solution
Sl.No | Option | Key Components | Remarks
1 | Using Phase Locked Loop ( PLL ) or frequency multiplier in combination with STM32H7 or ESP32-S3 | Voltage Controlled Oscillator | 511-STM32H725AEI6 PIC24FJ128GC010 FAMILY(MicroChip)
 |  | LUT ( Look Up Table ) | 
2 | Using FPGA | FPGA | 
 |  | Numerically Controller Oscillator (NCO) / LUT | 
3 | Using Programmable Oscillator module with Microcontroller | Programmable Oscillator Module | 
 |  | Microcontroller | 
4 | Using Direct Digital Synthesis ( DDS ) chip | AD 9833 | 
 |  | AD 9850 | 

## Slide 9
Development of Solution
Sl No | MCU Type | Processor speed | AHB Speed | APB1 Speed | APB2 Speed | Flash Memory | SRAM | Package | Price | Remarks
1 | STM32F401RE | 84 MHz | 84 MHz | 42 MHz | 42 MHz | 512 KB | 96 KB | LQFP64 | Rs.585 | 
2 | STM32F446RE | 180 MHz | 180 MHz | 45 MHz | 90 MHz | 512 KB | 112 KB | LQFP64 | Rs.783 | 
3 | STM32F722RE | 216 MHz | 216 MHz | 54 MHz | 108 MHz | 512 KB | 256 KB | LQFP64 | Rs.938 | 
4 | STM32H7A3A | 280 MHz | 280 MHz | 140 MHz | 140 MHz | 1024 KB | 128 KB | LQFP64 | Rs.1202 | 
5 | STM32H725ZG | 550 MHz | 275 MHz | 275 MHz | 275 MHz | 1024 KB | 128 KB | LQFP144 | Rs.1017 | 
6 | STM32H725RG | 550 MHz | 275 MHz | 138 MHz | 138 MHz | 1024 KB | 128 KB | VFQFPN68 | Rs.1035 | 
7 | STM32H725VG | 550 MHz | 275 MHz | 275 MHz | 275 MHz | 1024 KB | 128 KB | LQFP100 | Rs.1009 | 

## Slide 11
Sine Wave
Square Wave
Triangle Wave
Amplitude :
Frequency :
Phase Shift :
Sampling rate :
Output Channel :
Repetition :
User Interface Screen
Sine Wave
Square Wave
Triangle Wave
Amplitude :
Frequency :
Phase Shift :
Sampling rate :
Output Channel :
Repetition :

## Slide 13
STM32F446 family microcontrollers
Pending DAC Details :
Full scale output
Noise, linearity, resolution, tolerance, max resolution

## Slide 14
Digital To Analog Converters
Parallel
interface
Parallel &
buffered
QSPI with
RAM
SPI w/o
RAM
SPI with
RAM
>200 MSPS
120 to 180
MSPS
80 to 100
MSPS
12 bit resolution
AD9744
12 bit resolution
AD9706
AD9750
12 bit resolution
12 bit resolution
AD9753
10 bit resolution
AD9751
12 bit resolution
10 bit resolution
12 bit resolution
AD9102
08 bit resolution
AD9709
12 bit resolution
(Rs.2400)
(Rs.5600)
(Rs.1580)
(Rs.5210)
(Rs.1680)
(Rs.2300)
(100 MHz)
(60 MHz)
(100 MHz)
(60 MHz)
(60 MHz)
NIL
(Rs.1900)
(50 MHz)
12 bit resolution
NIL
10 bit resolution
08 bit resolution
NIL
NIL
AD9763
AD9765
(Rs.3370)
(Rs.2245)
AD7245A
AD3551R
AD9837
AD9833

## Slide 15
PC
SD Card
Touch
Screen
Driver
User Input Section
Level 1 Block Diagram of AWG (22-01-2024)
Clock Input
STM32 F401
Flash/
RAM
GPIO Ports
MMC
SPI
ADC
USB
VDD
Clock
In/Out
RTC
GPIO
Crystal
oscillator
Clock Output
AD9744
Clock
Input
Digital Data Input
VDD
VREF
IOUT
Power
Supply
(DC Adaptor)
230 V AC
Resistor
Output
Amplifier
Output
Wave
Clock Output ( 1 – 10 MHz )
(I to V)
Terminal Application
84MHz
3.3V
3.3
3.3V
Voltage Regulator
84 MHz ( Max )

## Slide 16
PC
Terminal
Application
SD Card
Touch
Screen
Driver
User
Input Section
Level 1 Block Diagram of AWG (29-01-2024)
Clock Input
STM32H723VG
Cortex M7
1MB Flash
564 kB SRAM
LTDC
2 X MMC
2 X Octo-SPI
USB
VDD
Clock
In
GPIO Port
( 220 MHz)
Oscillators
32.768 kHz LSE
Clock Output
AD9744
Clock
Input
Digital Data Input
AVDD
VREF
IOUTA
DC Power
connector
230 V AC
50Ω
Output
Amplifier
Clock Output ( 1 – 10 MHz )
550 MHz
LDO
Voltage
Regulator
<=10pF lines
Clock
Out
Type C
USB
Receptacle
5V
5V
3.3V
3.3V
LPF
CD1
CD2
Notes:
CD : Decoupling capacitor
LPF : Low Pass Filter
RL : Resistance
1.2V
LDO
IOUTB
50Ω
VoutA
VoutB
(Optional)
+
-
DVDD
FS Adj
50Ω
Diff
Amp
Oscillators
4 – 50 MHz HSE
SD Card Breakout
AHB4 275 MHz
64 MHz HSI
48 MHz HSI
4 MHz CSI
32 KHz LSI
CLK
CMD
D0 ( 3MBps)
Card Detect
25 MHz
D+
D-
BNC2
BNC1
5V
UART
GND
Rx
Tx
AC Adapter
External area
Signal Conditioning
DMA
12-bit parallel synchronous transmission
IDMA

## Slide 17
PC
Terminal
Application
SD Card
Touch
Screen
Driver
User
Input Section
Level 1 Block Diagram of AWG (05-02-2024)
Clock Input
STM32H723VG
Cortex M7
1MB Flash
564 kB SRAM
LTDC
2 X MMC
2 X Octo-SPI
USB
VDD
Clock
In
GPIO Port
( 220 MHz)
Clock Output
AD9744
Clock
Input
Digital Data Input
AVDD
VREF
IOUTA
50Ω
Output
Amplifier
Clock Output ( 1 – 10 MHz )
550 MHz
<=10pF lines
Clock
Out
5V
5V
3.3V
3.3V
Notes:
CD : Decoupling capacitor
LPF : Low Pass Filter
RL : Load Resistance
IOUTB
50Ω
VoutA
VoutB
(Optional)
+
-
DVDD
FS Adj
50Ω
Diff
Amp
AHB4 275 MHz
64 MHz HSI
48 MHz HSI
4 MHz CSI
32 KHz LSI
SD Card Breakout
CLK
CMD
D0 ( 3MBps)
Card Detect
25 MHz
D+
D-
Type C
USB
Receptacle
Oscillators
32.768 kHz LSE
LDO
Voltage
Regulator
Oscillators
4 – 50 MHz HSE
BNC2
BNC1
5V
UART
GND
Rx
Tx
DC Power
connector
230 V AC
AC Adapter
External area
Signal Conditioning
DMA
12-bit parallel synchronous transmission
IDMA
DAC
BNC3
LPF
CD1
CD2
3.3V

## Slide 18
STM32F446VET6
Cortex M4
512kB Flash
112 kB SRAM
GPIO
USB
VDD
Clock
In
SPI 45 MHz)
180 MHz
Clock
Out
AHB1 180 MHz
64 MHz HSI
48 MHz HSI
4 MHz CSI
32 KHz LSI
25 MHz
UART
DMA
DAC
SPI
AD9102
CLKVDD
SPI Interface
AVDD
VREF
IOUTP
PC
Terminal
Application
SD Card
Touch
Screen
Driver
User
Input Section
Type C
USB
Receptacle
SD Card Breakout
CLK
CMD
D0 ( 3MBps)
Card Detect
GND
Rx
Tx
DC Power
connector
230 V AC
AC Adapter
External area
Clock Output ( 1 – 10 MHz )
BNC1
Oscillators
32.768 kHz LSE
LDO
Voltage
Regulator
Oscillators
4 – 50 MHz HSE
BNC3
BNC2
5V
5V
5V
3.3V
IOUTN
+
-
Diff
Amp
DVDD
C
3.3V
FS Adj
SRAM
CS
SDIO
SDO
SCLK
RESET
CLKP
VoutA
VoutB
RL
RL
Notes:
CD : Decoupling capacitor
LPF : Low Pass Filter
RL : Load Resistance
1V
RF
Transformer
CLKP
C
C
Level 1 Block Diagram of AWG (12-02-2024)

## Slide 19
STM32F446RET6
Cortex M4
512kB Flash
112 kB SRAM
GPIO
USB
VDD
Clock
In
QSPI ( 90 MHz)
180 MHz
Clock
Out
AHB1 180 MHz
64 MHz HSI
48 MHz HSI
4 MHz CSI
32 KHz LSI
25 MHz
UART
DMA
DAC
SPI
AD9102
CLKVDD
SPI Interface
AVDD
VREF
IOUTP
PC
Terminal
Application
SD Card
Touch
Screen
Driver
User
Input Section
Type C
USB
Receptacle
SD Card Breakout
CLK
CMD
D0 ( 3MBps)
Card Detect
GND
Rx
Tx
DC Power
connector
230 V AC
AC Adapter
External area
Clock Output ( 1 – 10 MHz )
BNC1
Oscillators
32.768 kHz LSE
LDO
Voltage
Regulator
Oscillators
4 – 50 MHz HSE
BNC3
BNC2
5V
5V
5V
3.3V
IOUTN
+
-
Diff
Amp
DVDD
C
3.3V
FS Adj
SRAM
CS
SDIO
SDO
SCLK
RESET
CLKP
VoutA
VoutB
RL
RL
Notes:
CD : Decoupling capacitor
LPF : Low Pass Filter
RL : Load Resistance
1V
RF
Transformer
CLKP
C
C
Level 1 Block Diagram of AWG (12-02-2024)

## Slide 20
Evaluation Board

## Slide 21
Eval Board AD9102 – ARDZ – EBZ
Evaluation Board

## Slide 22
DAC
AD9102BCPZ
Logic Level Translator
(1.15V to 5.5 V)
EEPROM
On Board
Oscillator
Power Port : Arduino Conn 8pin
DIGI 0: Arduino Conn 8pin
DIGI 1: Arduino Conn 10pin
Eval Board AD9102 – ARDZ – EBZ – Block Diagram
Analog : Arduino Conn 6pin
Output 8 pin connector
Step Down
Regulator
DC-DC
Converter
Op Amp
Signal
Transformer
IOUTP1
IOUTN1
AMP_P1
AMP_N1
CKLN
CKLP
Signal
Transformer
Ext Clock Input
Jumper
SPI Interface
UART
SPI_SCK
SPI_MISO
SPI_MOSI
SPI_CS
TRIGGER
RESET
SDA
I2C
SCL
R
R
R
SPI_CS
TRIGGER
RESET
SPI_SCK
SPI_MISO
SPI_MOSI
6 pin connector
DC Power
Jack
DVDD
CLKVDD
Linear
Reg
3.6V
AVDD1
Linear
Reg
Power Input
CVDDX
CVDDX
3.6V
5.2V
DVDD
Output
Jumper
Linear
Reg

## Slide 23
Eval Board AD9102 – ARDZ – EBZ + SDP-K1